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Description:
This guide will assist with the installation of your replacement Watchdog O2 Puck.

Items:
· New Oxygen Puck
· Something to fan air over the puck

Instructions:

1.  First open up the right door of the unit and locate the white colored Watchdog Oxygen Sensor box.
2. Undo the clasps securing the lid of the box.
3. There is a circular screw on well that holds the 02 puck.  Unscrew it and remove the old one.
4. Following the instructions that come inside tbhe box with your new puck, place the new one in the well with the plastic pull tab pointing outwards.
5. Remove metal shorting clip from sensor.
6. Immediately fan air over the sensor until the display on the left side of the watchdog displays 20.9.  This may take several minutes.  If the display isn’t getting all the way to 20.9 after a few minutes of fanning air into it, hit the span button until you reach 20.9.  
7. Replace the screw on lid  and hand tighten.
8. Check the rotometer inside the watchdog, turn the small knob open or closed to target a flow rate near 0.8-1.0 SCFM.
9. Secure the clasps on the outside of the Watchdog box.
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